## Slide 1: Self-attenuation in the low-energy range: an experimental study on 210Pb
Self-attenuation effect:
The number of exiting photons is lower than the number of emitted photons
Depends on the source-detector geometry and photons energy
Motivation: Quantitative analysis of environmental samples generally involves volume source with low radioactivity and an unknown composition of the sample matrix. Unless an efficiency calibration with the same geometry and matrix is available, it is necessary to correct the raw counting rates for self-attenuation effects in order to obtain the correct sample activity.
This exercise focuses on the different steps required to calculate the final activity of unknown matrices, based on the calibration obtained with a reference matrix, in the same geometry. This includes: measurement of the linear attenuation coefficients, calculation of self-attenuations (between void and solid sample) and self attenuation corrective factors (between reference and unknown matrices).
M.-C. Lépy1, L. Chambon1, B. Sabot1, A. De Vismes2, R. Galea3, R. Idoeta4, P. Jodlowski5, K. Kanoutos6, K. Karfopoulos7, A. Meyer3, I. Mitsios6, V. Peyres8, P. Saganowski9, N. Salpadimos7, M.I. Savva10, O. Sima11, T.T. Thanh12, R. Townson4, Z. Tymiński9, T. Vasilopoulou10, L. Verheyen13, T. Vidmar13
1 CEA/LNHB, 2 IRSN, 3 NRC, 4 UPV/EHU, 5 AGH UST, 6 NTUA, 7 EEAE, 8 CIEMAT, 9 POLATOM, 10 NCSR “Demokritos”, 11 U. Bucharest, 12 VNUHCM, 13 SCK CEN
ON-GOING EXERCISE OF THE ICRM GAMMA SPECTROMETRY WORKING GROUP
Reference date: 1rst October 2022 – 12:00:00 UTC |  |  | 
 | Half-life (a) | Photon emission energy (keV) | Photon emission intensity (%)
210Pb | 22.23 (12) | 46.54 | 4.252 (40)
137Cs | 30.05 (8) | 661.66 | 84.99 (20)
4 batches circulating between 12 laboratories
For each matrix, one blank sample, one sample with 137Cs + 210Pb
At each participating laboratory,
Measurement and or calculation of linear attenuation coefficients
µ/r (cm-1)
137Cs acts as reference
(weak attenuation)
while 210Pb is a difficult case
(strong attenuation)
Detection efficiency calibration
for the two energies of interest
using the reference matrix
At LNHB, preparation of
3 samples with different densities
1 sample used as reference for calibration (activity known)
Epoxy polymer:
reference
Radionuclides
(210Pb and 137Cs)
2 samples (A and B) with
unknown activity
Measurement of
the 2 unknown matrix samples (A and B)
Determination of the activities of 137Cs and 210Pb included in sample A and sample B
Calculation of the self-attenuation corrective factors (between the reference and unknown matrices)
(Efficiency transfer factors)
Preliminary results:
  Each participant had to provide a report with detailed measurement procedure:
    detector characteristics
    measurement position for the samples under study
    calibration geometry and reference efficiency (46.5 keV and 662.7 keV)
    peak shape (210Pb) – influence of scattered photons in left side of the peak
  Measurement of linear attenuation coefficients of the 3 materials
    Procedure and results for the 3 matrices (reference and samples A and B) and 2 energies (46.5 keV and 662.7 keV)
  Conclusion: Strong variations – Analysis of the different experimental approaches on going which show a huge interest of such study
210Pb | Ref | A | B
Relative mean value | 1.0 | 1.0 | 1.0
Relative standard deviation (%) | 18.5 | 7.2 | 35.6
137Cs | Ref | A | B
Relative mean value | 1.0 | 1.0 | 1.0
Relative standard deviation (%) | 8.8 | 8.3 | 7.2
  Next steps:
  After checkings of the first results, next steps will consist in:
    calculation of self-attenuation (matrix/vacuum) and self-attenuation corrective factors (unknown matrix/reference matrix), procedure (method, software, approximation, etc.), results for the 3 matrices 2 energies
  Final results: activity (210Pb and 137Cs) of samples A and B with associated uncertainty (budget)
  Final goal: give recommandations for gamma spectrometry with complex matrix
Chart: Relative attenuation coefficients for 137Cs (662 keV)
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Reference | 1.0491 | 0.97397 | 1.0272 | 1.0822 | 1.0996 | 1.052 | 0.82205 | 0.9157 | 0.97813
Matrix A | 1.1387 | 0.97527 | 0.97527 | 1.0652 | 0.91311 | 1.1084 | 0.95831 | 0.91591 | 0.94983
Matrix B | 1.1082 | 0.96314 | 0.9664 | 1.0586 | 0.88963 | 0.97869 | 0.9624 | 1.0957 | 0.97721
 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
 | 1.0884 | 1.0884 | 1.0884 | 1.0884 | 1.0884 | 1.0884 | 1.0884 | 1.0884 | 1.0884 | 1.0884
 | 0.91164 | 0.91164 | 0.91164 | 0.91164 | 0.91164 | 0.91164 | 0.91164 | 0.91164 | 0.91164 | 0.91164
 | 1.0832 | 1.0832 | 1.0832 | 1.0832 | 1.0832 | 1.0832 | 1.0832 | 1.0832 | 1.0832 | 1.0832
 | 0.91683 | 0.91683 | 0.91683 | 0.91683 | 0.91683 | 0.91683 | 0.91683 | 0.91683 | 0.91683 | 0.91683
 | 1.0719 | 1.0719 | 1.0719 | 1.0719 | 1.0719 | 1.0719 | 1.0719 | 1.0719 | 1.0719 | 1.0719
 | 0.9281 | 0.9281 | 0.9281 | 0.9281 | 0.9281 | 0.9281 | 0.9281 | 0.9281 | 0.9281 | 0.9281
Chart: Relative attenuation coefficients for 210Pb (46.54 keV)
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Reference | 1.0752 | 0.8234 | 1.3075 | 1.0468 | 0.82455 | 0.79029 | 1.2413 | 0.92684 | 0.96408
Matrix A | 1.0036 | 0.88438 | 1.0339 | 1.129 | 0.94485 | 1.0209 | 0.96241 | 1.0209
Matrix B | 1.0577 | 1.0145 | 1.0077 | 0.9546 | 1.0381 | 0.94006 | 0.9874
 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
 | 1.1848 | 1.1848 | 1.1848 | 1.1848 | 1.1848 | 1.1848 | 1.1848 | 1.1848 | 1.1848 | 1.1848
 | 0.8152 | 0.8152 | 0.8152 | 0.8152 | 0.8152 | 0.8152 | 0.8152 | 0.8152 | 0.8152 | 0.8152
 | 1.0722 | 1.0722 | 1.0722 | 1.0722 | 1.0722 | 1.0722 | 1.0722 | 1.0722 | 1.0722 | 1.0722
 | 0.92779 | 0.92779 | 0.92779 | 0.92779 | 0.92779 | 0.92779 | 0.92779 | 0.92779 | 0.92779 | 0.92779
 | 1.0425 | 1.0425 | 1.0425 | 1.0425 | 1.0425 | 1.0425 | 1.0425 | 1.0425 | 1.0425 | 1.0425
 | 0.95745 | 0.95745 | 0.95745 | 0.95745 | 0.95745 | 0.95745 | 0.95745 | 0.95745 | 0.95745 | 0.95745
+ Reference, + Matrix A, + Matrix B
Gamma detector (HPGe)
Solid sample
Void (Transparent) sample
γ
γ
First step and already very intersting results and matter to discuss